CHỈNH RS-232 CÂN GS

- Press “ZERO/TARE” key until display “PC HEC”

- TTL continuous output, 8 data bits, 1 stop bit.

- Baud rate:9600bps

- Data format:

(start 1)(start 2)(flag bit)(data 1)(date 2)(data 3)(decimals)(unit)(reserve 1)(reserve 2)( verification)(end 1)(end 2)

a): all data by Hexadecimal

b): start 1 2: FE 55

c): flag bit:

bit 7: 0: positive,  1:negative

bit 6:

bit5: 0: not in zero point,  1: in zero point

bit4: 0: stable,   1: unstable

bit3:

bit2:

bit1

bit0: Overload Flag

d): data 1 2 3 decimals places: for example: 5500------00 15 7c 00,  55.00-----00 15 7c 02

e): unit: g:00, ct:01……. Can be arbitrary layout, OxFF: Count mode

f): reserve 1 2:  00  00

g): verification: (flag bit)(data 1)(date 2)(data 3)(decimals)(unit)(reserve 1)(reserve 2)

h): end 1 2: FE AA